## Sheet1
Table A.1: Biocompatibility Evaluation Endpoints
Device categorization by |  |  | Biologic effect
Nature of body contact |  | Contact duration A - limited (≦ 24h) B - prolonged (> 24h to 30d) C - permanent (> 30d) | Cytotoxicity | Sensitization | Irritation or Intracutaneous reactivity | Systemic toxicity (acute) | Subchronic toxicity (subacute toxicity) | Genotoxicity | Implantation | Haemocompatibility
Category | Contact
Surface device | Intact skin | A | X | X | X
|  | B | X | X | X
|  | C | X | X | X
| Mucosal membrane | A | X | X | X
|  | B | X | X | X | O | O |  | O
|  | C | X | X | X | O | X | X | O
| Breached or compromised surface | A | X | X | X | O
|  | B | X | X | X | O | O |  | O
|  | C | X | X | X | O | X | X | O
External communicating device | Blood path, Indirect | A | X | X | X | X |  |  |  | X
|  | B | X | X | X | X | O |  |  | X
|  | C | X | X | O | X | X | X | O | X
| Tissue / bone / dentin | A | X | X | X | O
|  | B | X | X | X | X | X | X | X
|  | C | X | X | X | X | X | X | X
| Circulating blood | A | X | X | X | X |  | O |  | X
|  | B | X | X | X | X | X | X | X | X
|  | C | X | X | X | X | X | X | X | X
Implant device | Tissue / bone | A | X | X | X | O
|  | B | X | X | X | X | X | X | X
|  | C | X | X | X | X | X | X | X
| Blood | A | X | X | X | X | X |  | X | X
|  | B | X | X | X | X | X | X | X | X
|  | C | X | X | X | X | X | X | X | X
X = ISO 10993-1 : 2009 recommended endpoints for consideration
O = Additional FDA recommended endpoints for consideration